## Pre Cuise
Pre-Cruise Check Off Sheet
CTD System
X |  | Power Up
X |  | Data Verification at Deck Unit
X |  | Data Verification at Acquiring Computer
X |  | Pump Verification at Deck Unit
X |  | Rosette Conformation at Deck Unit
X |  | Rosette Psychical Rotation Verification
X |  | Navigation Verification at Deck Unit
X |  | Navigation Verification at Acquiring Computer
X |  | Calibration Coefficients Verification at Acquiring Computer
X |  | Time Date Verification at Acquiring Computer
X |  | Data from last cruise properly backed up to storage directory
X |  | Dry Test Cast
ADCP System
X |  | Power Up
X |  | Proper Transducer Verification
X |  | Transducer Coefficients Verification at Acquiring Computer
X |  | Data Verification at Acquiring Computer
X |  | Navigation Verification at Acquiring Computer
X |  | Time Date Verification at Acquiring Computer
X |  | Data from last cruise properly backed up to storage directory
X |  | Test Run
SMS System
X |  | Power Up
X |  | Calibration Coefficients Verification at Acquiring Computer
X |  | Data Verification at Acquiring Computer
X |  | Navigation Verification at Acquiring Computer
X |  | Time Date Verification at Acquiring Computer
X |  | Data from last cruise properly backed up to storage directory
X |  | Test Run
Logger Light
X |  | Power Up
X |  | Calibration Coefficients Verification at Acquiring Computer
X |  | Data Verification at Acquiring Computer
X |  | Time Date Verification at Acquiring Computer
X |  | Data from last cruise properly backed up to storage directory
X |  | Test Cast
N/A | XBT
Power Up
Calibration Coefficients Verification at Acquiring Computer
Data Verification at Acquiring Computer
Navigation Verification at Acquiring Computer
Time Date Verification at Acquiring Computer
Data from last cruise properly backed up to storage directory
Test Drop
Water Maker
X |  | Setting Verification
X |  | Production Verification
X |  | Tank Full
X |  | Spare Tank Full
X |  | Spare Filter Verification
N/A | LSC
Power Up
Count Verification
Notified Crew of Hot Work
Proper Warning in Place
N/A | Lab Incubator
Power Up
Operation Verification
Light Verification
Timer Verification
Coring Equipment
N/A |  | Box Core
Operational Part Verification
Function Verification
Gear Secured for Sea
N/A |  | Gravity Core
Operational Part Verification
Function Verification
Sufficient Core Liners & End Caps
Gear Secured for Sea
N/A |  | Smith Mac
Operational Part Verification
Function Verification
Gear Secured for Sea
Deck Gear
X |  | Deck Incubators 1 & 2
X |  |  | Operational Part Verification
X |  |  | Function Verification
X |  |  | Gear Secured for Sea
N/A |  | Pole Mount
Operational Part Verification
Function Verification
Gear Secured for Sea

## Instructions
DAILY CRUISEBOOK CHECKS
Naming the Document
The current name for this workbook is "cruisebook.xls". At the beginning of each cruise the workbook file will be copied to a new file with the cruise name and number as found on the schedule (i.e.1506)
Filling in the Document
CTD
On each of the "Day" sheets there is a section for the CTD. There are two check off boxes for the daily care of the CTD and the following sensor sections.
Daily Fluorometer
The first sensor section deals with the fluorometer, which is Voltage 0 in SeaSave. After the lens on the fluorometer has been cleaned a voltage signal check will be made. The zero reading is made by inserting the white cap on the fluorometer. This reading should be around 3.00 VDC. If the zero reading is high the lens may require additional cleaning. A full scale reading is done by placing the black capon the fluorometer. This reading should be very close to 0.000 VDC
SMS
On each of the "Day" sheets there is a section for the SMS. There is a check off boxe for the daily care of the SMS and the following sensor sections. The time for each of the entries will be recorded from the SMS display.
Air Temperature
In the "Air Temperature" section an entry for dry air temperature from the display of the SMS will be entered. The "Aux Sensor Ship" is taken from the Dry air temperature reading from the Sling Psycrometer. Both of these reading will be entered in Degrees C.
Air Humidity
In the "Humidity" section an entry for air humidity from the display of the SMS will be entered. The "Aux Sensor Ship" is calculated from the air humidity reading from the Sling Psycrometer. Both of these readings will be entered as %RH.
Air Barometer
In the "Barometer" section an entry for barometer from the display of the SMS will be entered. The "Aux Sensor Ship" is taken from the barometer reading abtainned from the ship barometer located on the bridge. Both of these reading will be entered in mili-bars.
Water Temperature
In the "Water Temperature" section an entry for water temperature from the display of the SMS will be entered. The "Aux Sensor Ship" is taken from the water temperature reading abtainned from the ship water temperature located on the bridge. Both of these reading will be entered a degree C.
ADCP
On each of the "Day" sheets there is a section for the ADCP. This section contains entries for a reasonable check for both SOG and COG.
Speed over Ground (SOG)
In the "SOG" section an entry for ADCP SOG taken from the Nav portion on the ADCP. Look at the Mag and Made Good box which is displayed in m/s and will need to be converted to knots. The "Aux Sensor Ship" is taken from the SOG reading obtained from the ship DGPS located on the SMS display. Both of these reading will be entered in knots
Course over Ground (COG)
In the "COG" section an entry for ADCP COG taken from the Nav portion on the ADCP. Look at the Dir and Made Good box. The "Aux Sensor Ship" is taken from the COG reading obtained from the ship DGPS located on the SMS display. Both of these reading will be entered a kn/hour.
Salinity Samples
CTD&SMS at Sea
When a salinity sample is taken the reading are taken from the proper system. Record salinity, sample bottle number, and time and place them in the proper box. How to take a salinity sample is located in the procedure manual.
CTD&SMS in Lab
When ruining the samples in the lad the salinity is recorded with the proper sample bottle. Record these reading on the cruise summary sheet.
MilliQ Water Maker
Take the MilliQ handle and begin to dispense water into the sink. Watch the tank % and dispense atleast 5%. On the main screen you will see the numbers which should be recorded.
Using Forms
Form-Sta-01
This is the standard CTD Cast Sheet.
Form-Surv-01
This sheet is used when conducting surveys. The sheet can be copied and placed just in front of the instructions sheet in the workbook. Then it can be filled out as the cruise goes on. Additional sheets use as needed. This sheet does not need to be printed out.
Form-RS-01
This is the standard repair and service sheet. When this sheet is used it is copied and placed at the beginning of workbook. This sheet does not need to be printed out.
Form-Ship-01
This is the standard shipping sheet. When this sheet is used it is copied and placed at the beginning of workbook. This sheet can be printed off to provide a shipping label.

## Day 1 (NotCompleted)
CTD System |  | Fluorometer |  |  | Daily Maintenance
Blocked VDC |  | 2.5 |  |  | Rosette ( Fresh h2o rinse)
Clear VDC |  | 0.07 |  |  | Sensors and tubing ( Fresh h2o rinse)
Span Difference 0-5VDC |  | 2.43 |  |  |  |  | 1506
SMS System |  | SMS | Aux Sensor Ship | Sensor Difference | SMS Time of reading
Met
| Air Temp | 12.05 | 15.5556 | 3.5056 | 1917GMT
| Humidity | 84 | 86 | 2 | 1917GMT
| Barometer | 1027 | #NA | #VALUE! | 1917GMT
Water
| Temp | 10.79 | 10.08 | -0.71 | 1827GMT
ADCP System |  | SOG (Made Good) | COG (Made Good) | Comments
| ADCP | 0.618 | 201.3
| SMS | 0.7 | 192.1
| Difference | 0.082 | -9.2
Salinity Samples
|  | Sample Bottle # | Salinity | Sample Time | Comments
| CTD | #NA | #N/A
| SMS | #NA | #N/A
Water Maker |  | Feed Temp in C | Feed in us | Permeate in us | Rejection in % | Comments
|  | #N/A | #N/A | #N/A | #N/A
Daily Comments

## Day 2 (5-17)
CTD System |  | Fluorometer |  |  | Daily Maintenance
Blocked VDC |  | 4.22 |  |  | Rosette ( Fresh h2o rinse) | x
Clear VDC |  | 0.056 |  |  | Sensors and tubing ( Fresh h2o rinse) | x
Span Difference 0-5VDC |  | 4.164 |  |  |  |  | 1506
SMS System |  | SMS | Aux Sensor Ship | Sensor Difference | SMS Time of reading
Met
| Air Temp | 17.61 | 16.66 | -0.95 | 0745 GMT
| Humidity | 86 | 82 | -4 | 0745 GMT
| Barometer | 1021 | N/A | #VALUE! | 0745 GMT
Water
| Temp | 14.9271 | N/A | #VALUE! | 0746 GMT
ADCP System |  | SOG (Made Good) | COG (Made Good) | Comments
| ADCP | 5.7596 | 117.6
| SMS | 5.5 | 120.3
| Difference | -0.25961 | 2.7
Salinity Samples
|  | Sample Bottle # | Salinity | Sample Time | Comments
| CTD | #NA | #N/A
| SMS | #NA | #N/A
Water Maker |  | Feed Temp in C | Feed in us | Permeate in us | Rejection in % | Comments
|  | #N/A | #N/A | #N/A | #N/A
Daily Comments

## Day 3 (5-18)
CTD System |  | Fluorometer |  |  | Daily Maintenance
Blocked VDC |  | 0.83 |  |  | Rosette ( Fresh h2o rinse) | x
Clear VDC |  | 0.07 |  |  | Sensors and tubing ( Fresh h2o rinse) | x
Span Difference 0-5VDC |  | 0.76
SMS System |  | SMS | Aux Sensor Ship | Sensor Difference | SMS Time of reading
Met
| Air Temp | 19.6 | 20 | 0.4 | 16:30GMT
| Humidity | 80 | 80 | 0 | 16:30GMT
| Barometer | 1020 | N/A | #VALUE! | 16:30GMT
Water
| Temp | 16.53 | N/A | #VALUE! | 16:30GMT
ADCP System |  | SOG (Made Good) | COG (Made Good) | Comments
| ADCP | 6.84 | 121.2
| SMS | 6.7 | 119
| Difference | -0.14 | -2.2
Salinity Samples
|  | Sample Bottle # | Salinity | Sample Time | Comments
| CTD | #NA | #N/A
| SMS | #NA | #N/A
Water Maker |  | Feed Temp in C | Feed in us | Permeate in us | Rejection in % | Comments
|  | #N/A | #N/A | #N/A | #N/A
Daily Comments

## Day 4 (5-19)
CTD System |  | Fluorometer |  |  | Daily Maintenance
Blocked VDC |  | 0.86 | 0.46 with rag covering |  | Rosette ( Fresh h2o rinse) | x
Clear VDC |  | 0.05861 |  |  | Sensors and tubing ( Fresh h2o rinse) | x
Span Difference 0-5VDC |  | 0.80139
SMS System |  | SMS | Aux Sensor Ship | Sensor Difference | SMS Time of reading
Met
| Air Temp | 17.72 | 16.67 | -1.05 | 03:30GMT
| Humidity | 93 | 96 | 3 | 03:30GMT
| Barometer | 1010 | N/A | #VALUE! | 03:30GMT
Water
| Temp | 14.96 | N/A | #VALUE! | 03:12GMT
ADCP System |  | SOG (Made Good) | COG (Made Good) | Comments
| ADCP | 6.5469 | 58.8
| SMS | 6.7 | 56.9
| Difference | 0.15313 | -1.9
Salinity Samples
|  | Sample Bottle # | Salinity | Sample Time | Comments
| CTD | #NA | #N/A
| SMS | #NA | #N/A
Water Maker |  | Feed Temp in C | Feed in us | Permeate in us | Rejection in % | Comments
|  | #N/A | #N/A | #N/A | #N/A
Daily Comments

## Day 5 (5-20)
CTD System |  | Fluorometer |  |  | Daily Maintenance
Blocked VDC |  | 0.9 |  |  | Rosette ( Fresh h2o rinse) | x
Clear VDC |  | 0.0613 |  |  | Sensors and tubing ( Fresh h2o rinse) | x
Span Difference 0-5VDC |  | 0.8387
SMS System |  | SMS | Aux Sensor Ship | Sensor Difference | SMS Time of reading
Met
| Air Temp | 15.78 | 16.67 | 0.89 | 20:10GMT
| Humidity | 61 | 96 | 35 | 20:10GMT
| Barometer | 60 | N/A | #VALUE! | 20:10GMT
Water
| Temp | 14.6 | N/A | #VALUE! | 20:10GMT
ADCP System |  | SOG (Made Good) | COG (Made Good) | Comments
| ADCP | 6.5 | 38
| SMS | 6.6 | 39.8
| Difference | 0.1 | 1.8
Salinity Samples
|  | Sample Bottle # | Salinity | Sample Time | Comments
| CTD | #NA | #N/A
| SMS | #NA | #N/A
Water Maker |  | Feed Temp in C | Feed in us | Permeate in us | Rejection in % | Comments
|  | #N/A | #N/A | #N/A | #N/A
Daily Comments

## Day 6 (5-21)
CTD System |  | Fluorometer |  |  | Daily Maintenance
Blocked VDC |  | 3.7 |  |  | Rosette ( Fresh h2o rinse) | x
Clear VDC |  | 0.5 |  |  | Sensors and tubing ( Fresh h2o rinse) | x
Span Difference 0-5VDC |  | 3.2
SMS System |  | SMS | Aux Sensor Ship | Sensor Difference | SMS Time of reading
Met
| Air Temp | 11.83 | 10.5 | -1.33 | 21.13GMT
| Humidity | 80 | 90 | 10 | 21.13GMT
| Barometer | 1014 | N/A | #VALUE! | 21.13GMT
Water
| Temp | 13.33 | N/A | #VALUE! | 21.13GMT
ADCP System |  | SOG (Made Good) | COG (Made Good) | Comments
| ADCP | 6.5 | 105.4
| SMS | 6.5 | 106.2
| Difference | 0 | 0.8
Salinity Samples
|  | Sample Bottle # | Salinity | Sample Time | Comments
| CTD | #NA | #N/A
| SMS | #NA | #N/A
Water Maker |  | Permeate (Us): | MQ Res: | TOC(ppb): |  | Comments
|  | 10.4 | 18.2 | 2
Daily Comments

## Day 7 (5-22)
CTD System |  | Fluorometer |  |  | Daily Maintenance
Blocked VDC |  | 3.03 |  |  | Rosette ( Fresh h2o rinse) | x
Clear VDC |  | 0.4 |  |  | Sensors and tubing ( Fresh h2o rinse) | x
Span Difference 0-5VDC |  | 2.63
SMS System |  | SMS | Aux Sensor Ship | Sensor Difference | SMS Time of reading
Met
| Air Temp | 16.44 | 16 | -0.44 | 17:13GMT
| Humidity | 66 | 71 | 5 | 17:13GMT
| Barometer | 1014 | N/A | #VALUE! | 17.13GMT
Water
| Temp | 12.8 | N/A | #VALUE! | 17.13GMT
ADCP System |  | SOG (Made Good) | COG (Made Good) | Comments
| ADCP | 6.3 | 114.2
| SMS | 6.3 | 115
| Difference | 0 | 0.8
Salinity Samples
|  | Sample Bottle # | Salinity | Sample Time | Comments
| CTD | #NA | #N/A
| SMS | #NA | #N/A
Water Maker |  | Permeate (Us): | MQ Res: | TOC(ppb): |  | Comments
|  | 10.1 | 18.2 | 2
Daily Comments

## Day 8 (5-23)
CTD System |  | Fluorometer |  |  | Daily Maintenance
Blocked VDC |  | 0.038 |  |  | Rosette ( Fresh h2o rinse) | x
Clear VDC |  | 3.03 |  |  | Sensors and tubing ( Fresh h2o rinse) | x
Span Difference 0-5VDC |  | -2.992
SMS System |  | SMS | Aux Sensor Ship | Sensor Difference | SMS Time of reading
Met
| Air Temp | 16.05 | 13.9 | -2.15 | 17:34GMT
| Humidity | 36 | 60 | 24 | 17:34GMT
| Barometer | 1029 | N/A | #VALUE! | 17.34GMT
Water
| Temp | 9.87 | N/A | #VALUE! | 17.34GMT
ADCP System |  | SOG (Made Good) | COG (Made Good) | Comments
| ADCP | 6.79 | 56
| SMS | 7 | 55.9
| Difference | 0.21 | -0.1
Salinity Samples
|  | Sample Bottle # | Salinity | Sample Time | Comments
| CTD | #NA | #N/A
| SMS | #NA | #N/A
Water Maker |  | Permeate (Us): | MQ Res: | TOC(ppb): |  | Comments
|  | 9.9 | 18.2 | 2
Daily Comments

## Day 9 (5-24)
CTD System |  | Fluorometer |  |  | Daily Maintenance
Blocked VDC |  | 0.038 |  |  | Rosette ( Fresh h2o rinse) | x
Clear VDC |  | 3.03 |  |  | Sensors and tubing ( Fresh h2o rinse) | x
Span Difference 0-5VDC |  | -2.992
SMS System |  | SMS | Aux Sensor Ship | Sensor Difference | SMS Time of reading
Met
| Air Temp | 15.67 | 18.3 | 2.63 | 15:53GMT
| Humidity | 79 | 70 | -9 | 15:53GMT
| Barometer | 1029 | N/A | #VALUE! | 15:53GMT
Water
| Temp | 13.9 | N/A | #VALUE! | 15:53GMT
ADCP System |  | SOG (Made Good) | COG (Made Good) | Comments
| ADCP | 6.26 | 206.1
| SMS | 6.2 | 207.1
| Difference | -0.06 | 1
Salinity Samples
|  | Sample Bottle # | Salinity | Sample Time | Comments
| CTD | #NA | #N/A
| SMS | #NA | #N/A
Water Maker |  | Permeate (Us): | MQ Res: | TOC(ppb): |  | Comments
|  | 10.2 | 18.2 | 2
Daily Comments

## Cruise Summary
Cruise # | 1506 | P.I. | NK |  |  | Number of days filled out |  |  |  | 10
CTD System |  |  | Avg blocked 0VDC | STDev blocked VDC | Avg clear 5VDC |  | STDev clear VDC | Avg signal Span 0-5 VDC | Accuracy
Fluorometer |  |  | 1.7907 | 1.5934 | 0.80843 |  | 1.2703 | 0.98223
SMS System
Met |  |  | Air Temp in C |  | Humidity in % |  |  | Barometer in millibars
Sensors |  |  | Accuracy at 20 C +-0.3 C |  | Accuracy at 20 C 3% RH |  |  | Accuracy +- 11 milibar
|  |  | Avg Difference between SMS & Sling Air Temp |  | Avg Difference between SMS & Sling Humidity |  |  | Avg Difference between SMS & Ship Barometer
|  |  | #REF! |  | #REF! |  |  | #REF!
|  |  | STDev Difference between SMS & Sling |  | STDev Difference between SMS & Sling |  |  | STDev Difference between SMS & Ship
|  |  | #REF! |  | #REF! |  |  | #REF!
Water |  |  | Water Temp
Sensor |  |  | Accuracy
Avg Difference between SMS & Ship Water Temp in C
#REF!
ADCP System |  |  | Avg SOG in Knots. Difference between ADCP & Ship |  | Avg COG in Degrees. Difference between ADCP & Ship |  |  | We are unable to state any accuracy for the ADCP there are too many variables. We must rely on the experience for the technician to know what will be reasonable.
|  |  | #REF! |  | #REF!
|  |  | SDev SOG in Knots. Difference between ADCP & Ship |  | SDev COG in Degrees. Difference between ADCP & Ship
|  |  | #REF! |  | #REF!
Salinity Samples
CTD | Bottle # |  | #NA | #REF! | #REF! |  | #REF! | #REF! |  | Avg Difference
| Portasal Salinity |  | 0 | 0 | 0 |  | 0 | 0 |  | #REF!
| Salinity Difference |  | #NA | #REF! | #REF! |  | #REF! | #REF!
| Bottle # |  | #REF! | #REF! | #REF! |  | #REF! | #REF! |  | STDev
| Portasal Salinity |  | 0 | 0 | 0 |  | 0 | 0 |  | #REF!
| Salinity Difference |  | #REF! | #REF! | #REF! |  | #REF! | #REF!
SMS | Bottle # |  | #NA | #REF! | #REF! |  | #REF! | #REF! |  | Avg Difference
| Portasal Salinity |  | 0 | 0 | 0 |  | 0 | 0 |  | #REF!
| Salinity Difference |  | #NA | #REF! | #REF! |  | #REF! | #REF!
| Bottle # |  | #REF! | #REF! | #REF! |  | #REF! | #REF! |  | STDev
| Portasal Salinity |  | 0 | 0 | 0 |  | 0 | 0 |  | #REF!
| Salinity Difference |  | #REF! | #REF! | #REF! |  | #REF! | #REF!
Water Maker |  |  | Avg Temp in C | Avg Feed in us | Avg Permeate in us |  | Avg Rejection
|  | AVG | #REF! | #REF! | #REF! |  | #REF!
|  | STDev | #REF! | #REF! | #REF! |  | #REF!

## Form-Sta-01
R.V. HUGH R. SHARP |  | CRUISE Name | OPERATOR Name | Chief Scientist |  | DATE | Station | Cast
Lat
|  |  | CTD |  |  |  | Long
Ramp |  | Niskin # | Depth (m) | Temp C | Salinity | Sample Bottle # | Remarks
Position 1
Position 2
Position 3
Position 4
Position 5
Position 6
Position 7
Position 8
Position 9
Position 10
Position 11
Position 12
Sensor Measurements |  |  | Temp | Salinity | Fluorometer | /VOLT 0 | Cable #
| Time LMT | SAIL |  |  |  |  | Max. Out
| Time GMT | CTD |  |  |  |  | Max. Tension
Deck Unit S/N |  | CTD Con File |  | Depth CTD |  |  | Salinity Samples
403 |  | 2012-C.CON |  | Depth SMS |  |  | yes _____ | no _____
Comments
LOGERLIGHT
Binary File Name |  |  | ASCII File Name |  |  | Depth of Profile
Comments
Sediment
Core # |  | Sampler |  | Penetration and Description
Comments
Tows
Type of Tow |  |  |  | Start Time |  |  | Finish Time
Comments
Mooring
Start Time | [5 empty] | Finish Time
Description

## Form-Surv-01
R.V. Hugh R Sharp |  | CRUISE Name | OPERATOR Name |  |  | Chief Scientist
|  |  | Position |  |  | Cable #
DATE | Time | Cast | Station | Latitude | Longitude | Max. Out | Max. Tension | Comments

## Form-RS-01
Item |  | Date
Model # |  | Cruise
S/N |  | Operator
Repair | ____
Service | ____
If this is a repair explain what happened?
What was done to the item?
What parts need to be ordered for replacement spares?
If the science mission was interrupted explain.
Explain how we can improve.

## Form-Ship-01
Date
Shipped From:
University Of Delaware
700 Pilottown Rd.
Marine Operation Building Room 132
Lewes, Delaware 19958
Ship To:
Attn.
RMA #
Item Description:
Being returned for:
Shipper:
Tracking #